## Note
Created by | Yang Yang
Date | Feb 8_2016
Goal | Data archive
Table of Contents
Tab Name | Description
MDL | Describe the definition of method detection limit
CCVS | Describe the continuing calibration verification standards
Wood_foliage_results | Mercury concentration in wood and foliage of trees collected from Hubbard Brook
Wood_foliage_pool | Comparison of mercury pool in wood and in foliage
air-contamination result | The test of picking up indoor gas mercury for air-dried samples
calibration curve | absorbance vs. mercury content
raw data1 | Raw data for mercury concentration in wood and foliage
raw data2 | Raw data for air-contamination result

## MDL
Name | Weight (g) | Height | Hg [ng] | R [µg/kg] | True Value ug/Kg | Hg (ng) | CF
apple1 | 0.0054 | 0.0124 | 0.2356 | 43.6242 | 44 | 0.2376 | 19.1613
apple2 | 0.0054 | 0.0122 | 0.2314 | 42.8527 | 44 | 0.2376 | 19.4754
apple3 | 0.0047 | 0.0109 | 0.2039 | 43.3808 | 44 | 0.2068 | 18.9725
apple4 | 0.0049 | 0.011 | 0.2064 | 42.1244 | 44 | 0.2156 | 19.6
apple5 | 0.0052 | 0.0117 | 0.221 | 42.4981 | 44 | 0.2288 | 19.5556
apple6 | 0.0056 | 0.0122 | 0.2324 | 41.5016 | 44 | 0.2464 | 20.1967
apple7 | 0.0053 | 0.0117 | 0.221 | 41.6963 | 44 | 0.2332 | 19.9316
Average |  |  | 0.22167 | 42.5254 |  |  | 19.5562
Stdev |  |  | 0.012607 | 0.81069 |  |  | 0.41959
%RSD |  |  | 5.6875 | 1.9064 |  |  | 2.1456
MDL--98% |  |  | 0.039625 | 2.548 | 0.039625 | 2.548
T-TEST TABLE |  |  | 3.143 | 3.143
epa-procedure for mdl for 98% confid
Name | Weight (g) | Height | Hg [ng] | R [µg/kg]
blk | 0.2 | 0.0004 | 0.0078225 | 0.039112
blk | 0.2 | 0.0004 | 0.0078225 | 0.039112
blk | 0.2 | 0.0004 | 0.0078225 | 0.039112
blk | 0.2 | 0.0004 | 0.0078225 | 0.039112
blk | 0.2 | 0.0004 | 0.0078225 | 0.039112
blk | 0.2 | 0.0006 | 0.011734 | 0.058668
blk | 0.2 | 0.0006 | 0.011734 | 0.058668
Average(ng/g) |  | 0.00045714 | 0.00894 | 0.0447
|  | 9.759e-05 | 0.0019085 | 0.0095424
|  | 21.3478 | 21.3478 | 21.3478
LOD | 3 |  | 0.0057255 | 0.028627 | 0.060397 | 0.063537 | 0.079234
LOQ | 10 |  |  | 0.095424

## CCVs
Name | Hg [ng] | R [µg/kg] | Weight [g] | Area Peak | %R
160126dorm-primer | 21.6676 | 388.3084 | 0.0558 | 0.0214 | 94.7094
160126dorm-primer2 | 22.3333 | 406.0599 | 0.055 | 0.0221 | 99.039
160126dorm-primer3 | 20.9068 | 416.4706 | 0.0502 | 0.0206 | 101.5782
160126dorm-primer4 | 21.4774 | 406.0002 | 0.0529 | 0.0212 | 99.0244
2976ccv1 | 1.0604 | 57.6309 | 0.0184 | 0.052 | 94.4769
2976ccv2 | 0.9771 | 56.1547 | 0.0174 | 0.048 | 92.0569
2976ccv3 | 1.1562 | 61.8302 | 0.0187 | 0.0566 | 101.361
2976ccv4 | 0.7688 | 59.5969 | 0.0129 | 0.038 | 97.6998
2976ccv5 | 1.0229 | 64.3344 | 0.0159 | 0.0502 | 105.4662
2976ccv6 | 0.8105 | 63.8156 | 0.0127 | 0.04 | 104.6157
2976ccv7 | 0.8875 | 66.2333 | 0.0134 | 0.0437 | 108.5792
2976ccv8 | 0.9349 | 60.3142 | 0.0155 | 0.0443 | 98.8757
2976ccv9 | 1.1541 | 66.7134 | 0.0173 | 0.0565 | 109.3662
2976ccv10 | 0.7709 | 64.7801 | 0.0119 | 0.0381 | 106.1969
2976ccv11 | 0.7355 | 65.6672 | 0.0112 | 0.0364 | 107.6511
2976ccv12 | 0.6376 | 60.7215 | 0.0105 | 0.0317 | 99.5434
2976ccv13 | 0.8875 | 60.7895 | 0.0146 | 0.0437 | 99.6549
2976ccv14 | 0.8688 | 61.1817 | 0.0142 | 0.0428 | 100.2979
2976ccv15 | 0.7875 | 63.0037 | 0.0125 | 0.0389 | 103.2848
2976ccv16 | 0.8771 | 63.5588 | 0.0138 | 0.0432 | 104.1948
160126apple-qcs | 0.2439 | 43.554 | 0.0056 | 0.0128 | 98.9864
160126-apple-qcs2 | 0.2272 | 42.0813 | 0.0054 | 0.012 | 95.6393
160126apple-qcs3 | 0.2269 | 45.379 | 0.005 | 0.012 | 103.1341
160126-apple-qcs4 | 0.1808 | 38.4643 | 0.0047 | 0.0098 | 87.4189
160126apple-qcs5 | 0.2147 | 37.6739 | 0.0057 | 0.0114 | 85.6225
160126-apple-qcs6 | 0.2356 | 48.0756 | 0.0049 | 0.0124 | 109.2627
160126-ccb1 | 0 | 0 | 0 | 0.0001
160126-ccb2 | 0 | 0 | 0 | 0.0007
160126-ccb3 | 0 | 0 | 0 | 0.0006
160126-ccb4 | 0 | 0 | 0 | 0.0003
160126-ccb5 | 0 | 0 | 0 | 0.0006
160126-ccb6 | 0 | 0 | 0 | 0.0007
160126-ccb7 | 0 | 0 | 0 | 0.0007
160126-ccb8 | 0 | 0 | 0 | 0.0002
160126-ccb9 | 0 | 0 | 0 | 0.0003

## wood_foliage_results
Name | Concentration (ng/g)
sp123l | 16.7994 | 17.2924 |  | Leaves |  | Wood
sp123l | 17.7854 |  | Spruce | 17.2924 |  | 2.0654
sp456l | 17.565 | 17.4434 |  | 17.4434 |  | 1.9614
sp456l | 17.3219 |  |  | 23.826 | 19.5206 | 1.7965 | 1.9411
sp789l | 24.1623 | 23.826 |  |  | 2.1531 |  | 0.078314
sp789l | 23.4898 |  | Fir
fir123l | 28.6396 | 28.7458 |  | 28.7458 |  | 1.5544
fir123l | 28.8519 |  |  | 39.9999 |  | 1.1426
fir456l | 40.2001 | 39.9999 |  | 30.8944 | 33.2134 | 1.1871 | 1.2947
fir456l | 39.7997 |  |  |  | 3.4495 |  | 0.13049
fir789l | 30.7727 | 30.8944 | AB | 25.7246 |  | 1.2215
fir789l | 31.0162 |  |  | 20.6598 |  | 2.2719
ab123l | 25.7275 | 25.7246 |  | 21.6221 | 22.6688 | 1.6953 | 1.7296
ab123l | 25.7216 |  |  |  | 1.5529 |  | 0.30371
ab456l | 20.7406 | 20.6598 | YB | 17.5429 |  | 1.8945
ab456l | 20.5791 |  |  | 16.9168 |  | 2.8396
ab789l | 21.6004 | 21.6221 |  | 17.9453 | 17.4683 | 3.5189 | 2.751
ab789l | 21.6437 |  |  |  | 0.29922 |  | 0.47098
yb123l | 17.4418 | 17.5429 | SM | 16.5133 |  | 1.434
yb123l | 17.644 |  |  | 16.4787 |  | 0.4108
yb456l | 16.8663 | 16.9168 |  | 20.9441 | 17.9787 | 0.5969 | 0.8139
yb456l | 16.9673 |  |  |  | 1.4827 |  | 0.31467
yb789l | 17.8793 | 17.9453
yb789l | 18.0112 |  | Other hardwoods |  | 19.372 |  | 1.7648
sm123l | 16.5165 | 16.5133 |  |  | 1.655 |  | 0.55948
sm123l | 16.5101 |  | Other conifers |  | 26.367 |  | 1.6179
sm456l | 16.794 | 16.4787 |  |  | 6.8464 |  | 0.32321
sm456l | 16.1635
sm789l | 21.1846 | 20.9441
sm789l | 20.7036
sp123w | 2.3322 | 2.0654
sp123w | 1.7987
sp456w | 1.9707 | 1.9614
sp456w | 1.9521
sp789w | 1.9938 | 1.7965
sp789w | 1.5991
fir123w | 1.6009 | 1.5544
fir123w | 1.5079
fir456w | 1.1831 | 1.1426
fir456w | 1.1021
fir789w | 1.2733 | 1.1871
fir789w | 1.1008
ab123w | 1.3571 | 1.2215
ab123w | 1.0859
ab456w | 2.2663 | 2.2719
ab456w | 2.2775
ab789w | 1.5881 | 1.6953
ab789w | 1.8025
sm123w | 1.7561 | 1.434
sm123w | 1.1119
sm456w | 0.4286 | 0.4108
sm456w | 0.393
sm789w | 0.5028 | 0.5969
sm789w | 0.691
yb123w | 1.8408 | 1.8945
yb123w | 1.9483
yb456w | 2.9642 | 2.8396
yb456w | 2.7151
yb789w | 3.4079 | 3.5189
yb789w | 3.6298

## wood_foliage_pool
Hg concentration |  |  |  |  | Biomass of wood and foliage in 2007
|  | Hg concentration (ng g-1) |  |  | (http://www.hubbardbrook.org/w6_tour/biomass-stop/biomassw6.htm)
| Species | Wood | Foliage
| American beech | 1.7296 | 22.6688 |  |  |  | Biomass (Mg ha-1)
| Sugar maple | 0.8139 | 17.9787 |  |  | Species | Wood | Foliage
| Yellow birch | 2.751 | 17.4683 |  |  | AB | 67.98 | 2.26 | 1.582 | Sugar maple | 39.67 | 1.21 |  | Sugar maple | 21.23 | 0.85 | Sugar maple | 3.6 | 0.16
| Balsam fir | 1.2947 | 33.2134 |  |  | SM | 64.5 | 2.22 | 1.554 | Beech | 30.53 | 0.77 |  | Beech | 26.87 | 1.04 | Beech | 10.58 | 0.45
| Red spruce | 1.9411 | 19.5206 |  |  | BF | 8.36 | 0.18 | 0.126 | Yellow birch | 21.76 | 0.65 |  | Yellow birch | 16.59 | 0.79 | Yellow birch | 8.86 | 0.54
| Other hardwoods | 1.7648 | 19.372 |  |  | RS | 8 | 0.17 | 0.119 | White ash | 0.23 | 0.01 |  | Mountain maple | 0.79 | 0.05 | Mountain maple | 0.9 | 0.07
| Other conifers | 1.6179 | 26.367 |  |  | YB | 47.21 | 1.98 | 1.386 | Mountain maple | 0.26 | 0.01 |  | Striped maple | 0.24 | 0.01 | Striped maple | 2.21 | 0.11
[6 empty] | WA | 0.23 | 0.01 | 0.007 | Striped maple | 0.41 | 0.02 |  | Pin cherry | 0.41 | 0.03 | Pin cherry | 1.08 | 0.09
[6 empty] | STM | 2.86 | 0.14 | 0.098 | Pin cherry | 0.53 | 0.03 |  | Balsam fir | 0.02 | 0 | Choke cherry | 0.04 | 0
[6 empty] | PC | 2.02 | 0.15 | 0.105 | Red spruce | 0.34 | 0.01 |  | Red spruce | 1.55 | 0.03 | Balsam fir | 8.34 | 0.18
[6 empty] | CHOKE CHERRY | 0.04 | 0 | 0 |  |  |  |  | White birch | 1.71 | 0.11 | Red spruce | 6.11 | 0.13
[6 empty] | WB | 13.48 | 0.9 | 0.63 |  |  |  |  | Mountain ash | 0.02 | 0 | White birch | 11.77 | 0.79
[6 empty] | MA | 0.68 | 0.04 | 0.028 | [7 empty] | Mountain ash | 0.66 | 0.04
[6 empty] | RM | 1.95 | 0.09 | 0.063 | [7 empty] | Red maple | 1.95 | 0.09
Calculation of Hg pool
Biomass (Mg ha-1)
| Species | Wood | Foliage
| American beech | 67.98 | 1.59
| Sugar maple | 64.5 | 1.04
| Yellow birch | 47.2 | 0.53
| Balsam fir | 1.24 | 0.02
| Red spruce | 1.7 | 0.03
| Other hardwoods | 21.26 | 1.33
| Other conifers | 0 | 0
Concentrations
Wood |  |  |  |  | Foliage
| ng/g | mg/Mg | mg/ha | g/ha |  | ng/g | mg/Mg | mg/ha | g/ha
AB | 1.7296 | 1.7296 | 117.5759 | 0.11758 | AB | 22.6688 | 22.6688 | 36.0434 | 0.036043
SM | 0.8139 | 0.8139 | 52.4965 | 0.052497 | SM | 17.9787 | 17.9787 | 18.6979 | 0.018698
YB | 2.751 | 2.751 | 129.848 | 0.12985 | YB | 17.4683 | 17.4683 | 9.2582 | 0.0092582
BF | 1.2947 | 1.2947 | 1.6054 | 0.0016054 | BF | 33.2134 | 33.2134 | 0.66427 | 0.00066427
RS | 1.9411 | 1.9411 | 3.2999 | 0.0032999 | RS | 19.5206 | 19.5206 | 0.58562 | 0.00058562
other hardwood | 1.7648 | 1.7648 | 37.5202 | 0.03752 | other hardwood | 19.372 | 19.372 | 25.7647 | 0.025765
other conifer | 1.6179 | 1.6179 | 0 | 0 | other conifer | 26.367 | 26.367 | 0 | 0
Hit shift and reload to change the graphs on successive queries in some browsers
total hardwood mercury |  |  |  | 0.33744 | total hardwood mercury |  |  |  | 0.089764
total conifer mercury |  |  |  | 0.0049053 | total conifer mercury |  |  |  | 0.0012499
in total |  |  |  | 0.34235 | in total |  |  |  | 0.091014 | Graphs are for Watershed 6 1965, Lower hardwoods
All live trees
Stardnard error | [9 empty] | All trees >=2 cm
Wood |  |  |  |  | Foliage
| ng/g | mg/Mg | mg/ha | g/ha |  | ng/g | mg/Mg | mg/ha | g/ha | Software error:
AB | 0.30371 | 0.30371 | 20.6461 | 0.020646 | AB | 1.5529 | 1.5529 | 2.4691 | 0.0024691
SM | 0.31467 | 0.31467 | 20.2962 | 0.020296 | SM | 1.4827 | 1.4827 | 1.542 | 0.001542 | Can't | locate | object | method | "new" | via | package | "Chart::Bars" | (perhaps | you | forgot | to | load | "Chart::Bars"?) | at | /data/web/hbrsensor/cgi-bin/biomass/biomassw6.pl | line | 1232, | <CAT> | line | 3993
YB | 0.47098 | 0.47098 | 22.2305 | 0.02223 | YB | 0.29922 | 0.29922 | 0.15859 | 0.00015859
BF | 0.13049 | 0.13049 | 0.16181 | 0.00016181 | BF | 3.4495 | 3.4495 | 0.06899 | 6.899e-05 | For assistance, please contact Thomas Siccama
RS | 0.078314 | 0.078314 | 0.13313 | 0.00013313 | RS | 2.1531 | 2.1531 | 0.064594 | 6.4594e-05
other hardwood | 0.55948 | 0.55948 | 11.8944 | 0.011894 | other hardwood | 1.655 | 1.655 | 2.2012 | 0.0022012
other conifer | 0.32321 | 0.32321 | 0 | 0 | other conifer | 6.8464 | 6.8464 | 0 | 0

## air_contamination result
| Weigth [g] | Height | Hg [ng] | [ug/kg] | ave | se
Freeze dry | 0.0995 | 0.0023 | 0.0252 | 0.2532 | 0.24027 | 0.0064831
Freeze dry | 0.0992 | 0.0022 | 0.0231 | 0.233
Freeze dry | 0.0985 | 0.0022 | 0.0231 | 0.2346
Clean room | 0.0926 | 0.0023 | 0.0252 | 0.2721 | 0.26037 | 0.0096176
Clean room | 0.0941 | 0.0023 | 0.0252 | 0.2677
Clean room | 0.0958 | 0.0022 | 0.0231 | 0.2413
Cowetta soil lab | 0.0902 | 0.0029 | 0.0377 | 0.4179 | 0.39277 | 0.020981
Cowetta soil lab | 0.0955 | 0.0027 | 0.0335 | 0.3511
Cowetta soil lab | 0.0921 | 0.0029 | 0.0377 | 0.4093
Illick drying room | 0.0956 | 0.0038 | 0.0564 | 0.5904 | 0.70997 | 0.15962
Illick drying room | 0.0886 | 0.0055 | 0.0909 | 1.0261
Illick drying room | 0.0937 | 0.0034 | 0.0481 | 0.5134
Bartlett drying garage | 0.0878 | 0.0031 | 0.0419 | 0.4767 | 0.85917 | 0.22986
Bartlett drying garage | 0.0788 | 0.0059 | 0.1002 | 1.2713
Bartlett drying garage | 0.1057 | 0.0053 | 0.0877 | 0.8295
HB-Barn | 0.0967 | 0.0024 | 0.0273 | 0.2826 | 0.2604 | 0.012788
HB-Barn | 0.097 | 0.0022 | 0.0231 | 0.2383
HB-Barn | 0.0888 | 0.0022 | 0.0231 | 0.2603
HB-Storage room | 0.1129 | 0.0025 | 0.0294 | 0.2601 | 0.3763 | 0.093597
HB-Storage room | 0.1047 | 0.0026 | 0.0322 | 0.3073
HB-Storage room | 0.0968 | 0.0037 | 0.0544 | 0.5615
HB-Soil room | 0.1104 | 0.0022 | 0.0231 | 0.2093 | 0.23023 | 0.010467
HB-Soil room | 0.1024 | 0.0021 | 0.0201 | 0.2407
HB-Soil room | 0.0858 | 0.0015 | 0.0085 | 0.2407
Chamber contamination | 0.0973 | 0.2356 | 225.3681 | 2316.2185 | 2132.8534 | 111.3799
Chamber contamination | 0.0959 | 0.2155 | 206.2533 | 2150.7124
Chamber contamination | 0.1085 | 0.219 | 209.5818 | 1931.6292

## calibration curve
LIMS export of <yy160121.d80>
Version 1.0
Date 2016-01-22-15-51-33
Nr | Pos | SampleName | CreationDate | ProcessTime | Weigth [g] | ConcentrateWeight [g] | Type | State | CalFile | MethodFile | Height | Hg [ng] | [ug/kg] | DilutionWeight | FinalVolume | CalFactor | DilutionFactor | Statistics | Group | Remark
1 | 1 | blk | 42390860590 | 42390893391 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0029 | 0 | 0.0001 |  |  | 1
2 | 2 | blk | 42390860787 | 42390899293 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0 | 0 | 0.0001 |  |  | 1
3 | 3 | blk | 42390861400 | 42390905393 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0015 | 0 | 0.0001 |  |  | 1
4 | 4 | tuna | 42390861458 | 42390911631 | 0.0252 | 0.0252 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0744 | 71.82 | 2850 |  |  | 1
5 | 5 | tuna | 42390863287 | 42390917951 | 0.014 | 0.014 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0397 | 39.9 | 2850 |  |  | 1
6 | 6 | blk | 42390863356 | 42390924143 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0029 | 0 | 0.0001 |  |  | 1
7 | 7 | blk | 42390863391 | 42390930335 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0018 | 0 | 0.0001 |  |  | 1
8 | 8 | dorm4 | 42390863437 | 42390936655 | 0.0717 | 0.0717 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0298 | 29.397 | 410 |  |  | 1
9 | 9 | dorm4 | 42390863472 | 42390942997 | 0.0574 | 0.0574 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0237 | 23.534 | 410 |  |  | 1
10 | 10 | dorm4 | 42390863518 | 42390949201 | 0.0482 | 0.0482 | calibrate | finished | <data not saved> | yy150929.m80 | 0.785 | 19.762 | 410 |  |  | 1
11 | 11 | dorm4 | 42390863541 | 42390955405 | 0.0225 | 0.0225 | calibrate | finished | <data not saved> | yy150929.m80 | 0.4434 | 9.225 | 410 |  |  | 1
12 | 12 | blk | 42390863668 | 42390961597 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0021 | 0 | 0.0001 |  |  | 1
13 | 13 | blk | 42390863692 | 42390967777 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0022 | 0 | 0.0001 |  |  | 1
14 | 14 | 2976a | 42390863703 | 42390974108 | 0.0793 | 0.0793 | calibrate | finished | <data not saved> | yy150929.m80 | 0.2361 | 4.8373 | 61 |  |  | 1
15 | 15 | 2976a | 42390863784 | 42390980439 | 0.0493 | 0.0493 | calibrate | finished | <data not saved> | yy150929.m80 | 0.1435 | 3.0073 | 61 |  |  | 1
16 | 16 | 2976a | 42390863784 | 42390986631 | 0.0295 | 0.0295 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0886 | 1.7995 | 61 |  |  | 1
17 | 17 | 2976a | 42390863796 | 42390992789 | 0.0141 | 0.0141 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0382 | 0.8601 | 61 |  |  | 1
18 | 18 | blk | 42390863865 | 42390998969 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.001 | 0 | 0.0001 |  |  | 1
19 | 19 | blk | 42390863900 | 42391005138 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0009 | 0 | 0.0001 |  |  | 1
20 | 20 | applea | 42390863912 | 42391011446 | 0.0122 | 0.0122 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0273 | 0.5368 | 44 |  |  | 1
21 | 21 | applea | 42390863981 | 42391017766 | 0.0041 | 0.0041 | calibrate | finished | <data not saved> | yy150929.m80 | 0.008 | 0.1804 | 44 |  |  | 1
22 | 22 | applea | 42390864201 | 42391023958 | 0.0043 | 0.0043 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0086 | 0.1892 | 44 |  |  | 1
23 | 23 | applea | 42390864212 | 42391030127 | 0.0022 | 0.0022 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0051 | 0.0968 | 44 |  |  | 1
24 | 24 | applea | 42390864224 | 42391036261 | 0.0023 | 0.0023 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0072 | 0.1012 | 44 |  |  | 1
25 | 25 | applea | 42390864224 | 42391042384 | 0.0023 | 0.0023 | calibrate | finished | <data not saved> | yy150929.m80 | 0.011 | 0.1012 | 44 |  |  | 1
26 | 26 | applea | 42390864571 | 42391048530 | 0.0013 | 0.0013 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0051 | 0.0572 | 44 |  |  | 1
27 | 27 | applea | 42390864583 | 42391054664 | 0.0013 | 0.0013 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0043 | 0.0572 | 44 |  |  | 1
28 | 28 | blk | 42390864583 | 42391060798 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0007 | 0 | 0.0001 |  |  | 1
29 | 29 | blk | 42390864733 | 42391066921 | 0.0001 | 0.0001 | calibrate | finished | <data not saved> | yy150929.m80 | 0.0007 | 0 | 0.0001 |  |  | 1

## raw data1
LIMS export of <yy160205.d80>
Version 1.0
Date 2016-02-08-16-38-04
Nr | Pos | SampleName | CreationDate | ProcessTime | Weigth [g] | ConcentrateWeight [g] | Type | State | CalFile | MethodFile | Height | Hg [ng] | [ug/kg] | DilutionWeight | FinalVolume | CalFactor | DilutionFactor | Statistics | Group | Remark
1 | 1 | blk | 42405672673 | 42405682916 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0416 | 0.8438 | 0.8438 |  |  | 1
2 | 2 | blk | 42405672743 | 42405689178 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0024 | 0.0282 | 0.0282 |  |  | 1
3 | 3 | dorm-primer1 | 42405672743 | 42405695381 | 0.0586 | 0.0586 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0248 | 24.9009 | 424.9309 |  |  | 1
4 | 4 | dorm-primer2 | 42405672812 | 42405701585 | 0.0505 | 0.0505 | normal | finished | yy160121.c80 | yy150929.m80 | 0.8455 | 17.5887 | 348.2905 |  |  | 1
5 | 5 | blk | 42405672824 | 42405707789 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0032 | 0.0439 | 0.0439 |  |  | 1
6 | 6 | 2976ccv1 | 42405672847 | 42405713981 | 0.0182 | 0.0182 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0619 | 1.2673 | 69.631 |  |  | 1
7 | 7 | 2976ccv2 | 42405672893 | 42405720196 | 0.0129 | 0.0129 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0436 | 0.8854 | 68.639 |  |  | 1
8 | 8 | ccb1 | 42405672905 | 42405726400 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0012 | 0.0023 | 0.0023 |  |  | 1
9 | 9 | apple-qcs1 | 42405672962 | 42405732604 | 0.005 | 0.005 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0135 | 0.2585 | 51.6966 |  |  | 1
10 | 10 | apple-qcs2 | 42405673032 | 42405738854 | 0.0049 | 0.0049 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0141 | 0.2717 | 55.4524 |  |  | 1
11 | 11 | method-blk | 42405673032 | 42405745057 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0015 | 0.0085 | 0.0085 |  |  | 1
12 | 12 | blk | 42405739398 | 42405751250 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0011 | 0.0005 | 0.0005 |  |  | 1
13 | 13 | blk | 42405739421 | 42405757453 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0009 | 0 | 0 |  |  | 1
14 | 14 | blk | 42405739444 | 42405763703 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
15 | 15 | blk | 42405752442 | 42405769965 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.001 | 0 | 0 |  |  | 1
16 | 16 | blk | 42405752453 | 42405776226 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0009 | 0 | 0 |  |  | 1
17 | 17 | 2976ccv3 | 42405786574 | 42405788263 | 0.0173 | 0.0173 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0565 | 1.1541 | 66.7134 |  |  | 1
18 | 18 | 2976ccv4 | 42405786631 | 42405794513 | 0.0119 | 0.0119 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0381 | 0.7709 | 64.7801 |  |  | 1
19 | 19 | blk | 42405790740 | 42405800717 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0013 | 0.0052 | 0.0052 |  |  | 1
20 | 20 | blk | 42405790775 | 42405806921 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0009 | 0 | 0 |  |  | 1
21 | 21 | apple-qcs3 | 42405790787 | 42405813171 | 0.0052 | 0.0052 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0155 | 0.2993 | 57.5593 |  |  | 1
22 | 22 | apple-qcs4 | 42405790879 | 42405819386 | 0.005 | 0.005 | normal | finished | yy160121.c80 | yy150929.m80 | 0.012 | 0.2269 | 45.379 |  |  | 1
23 | 23 | blk | 42405792754 | 42405825590 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.001 | 0 | 0 |  |  | 1
24 | 24 | blk | 42405834039 | 42405834074 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0012 | 0.0018 | 0.0018 |  |  | 1
25 | 25 | blk | 42405842685 | 42405922199 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0021 | 0.0218 | 0.0218 |  |  | 1
26 | 26 | blk | 42405842789 | 42405928449 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
27 | 27 | sp123l | 42405842789 | 42405934652 | 0.0885 | 0.0885 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0725 | 1.4867 | 16.7994 |  |  | 1
28 | 28 | sp123l | 42405842858 | 42405940856 | 0.0871 | 0.0871 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0755 | 1.5491 | 17.7854 |  |  | 1
29 | 29 | sp456l | 42405842870 | 42405947060 | 0.0901 | 0.0901 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0771 | 1.5826 | 17.565 |  |  | 1
30 | 30 | sp456l | 42405842905 | 42405953263 | 0.0926 | 0.0926 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0781 | 1.604 | 17.3219 |  |  | 1
31 | 31 | sp789l | 42405842916 | 42405959467 | 0.0888 | 0.0888 | normal | finished | yy160121.c80 | yy150929.m80 | 0.1041 | 2.1456 | 24.1623 |  |  | 1
32 | 32 | sp789l | 42405842986 | 42405965659 | 0.0903 | 0.0903 | normal | finished | yy160121.c80 | yy150929.m80 | 0.1029 | 2.1211 | 23.4898 |  |  | 1
33 | 33 | fir123l | 42405842986 | 42405971863 | 0.0899 | 0.0899 | normal | finished | yy160121.c80 | yy150929.m80 | 0.1247 | 2.5747 | 28.6396 |  |  | 1
34 | 34 | fir123l | 42405843032 | 42405978067 | 0.0891 | 0.0891 | normal | finished | yy160121.c80 | yy150929.m80 | 0.1245 | 2.5707 | 28.8519 |  |  | 1
35 | 35 | fir456l | 42405843043 | 42405984270 | 0.0888 | 0.0888 | normal | finished | yy160121.c80 | yy150929.m80 | 0.1725 | 3.5698 | 40.2001 |  |  | 1
36 | 36 | fir456l | 42405843090 | 42405990474 | 0.091 | 0.091 | normal | finished | yy160121.c80 | yy150929.m80 | 0.175 | 3.6218 | 39.7997 |  |  | 1
37 | 37 | fir789l | 42405843125 | 42405996666 | 0.0895 | 0.0895 | normal | finished | yy160121.c80 | yy150929.m80 | 0.1333 | 2.7542 | 30.7727 |  |  | 1
38 | 38 | fir789l | 42405843159 | 42406002870 | 0.0906 | 0.0906 | normal | finished | yy160121.c80 | yy150929.m80 | 0.136 | 2.8101 | 31.0162 |  |  | 1
39 | 39 | ab123l | 42405843159 | 42406009074 | 0.0916 | 0.0916 | normal | finished | yy160121.c80 | yy150929.m80 | 0.1142 | 2.3566 | 25.7275 |  |  | 1
40 | 40 | ab123l | 42405878206 | 42406015277 | 0.0903 | 0.0903 | normal | finished | yy160121.c80 | yy150929.m80 | 0.1126 | 2.3227 | 25.7216 |  |  | 1
41 | 1 | ab456l | 42405882523 | 42406021469 | 0.0917 | 0.0917 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0924 | 1.9019 | 20.7406 |  |  | 1
42 | 2 | ab456l | 42405882569 | 42406027673 | 0.0908 | 0.0908 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0908 | 1.8686 | 20.5791 |  |  | 1
43 | 3 | ab789l | 42405888495 | 42406033877 | 0.0894 | 0.0894 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0938 | 1.9311 | 21.6004 |  |  | 1
44 | 4 | ab789l | 42405888530 | 42406040081 | 0.0872 | 0.0872 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0917 | 1.8873 | 21.6437 |  |  | 1
45 | 5 | 2976ccv5 | 42405888553 | 42406046273 | 0.0112 | 0.0112 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0364 | 0.7355 | 65.6672 |  |  | 1
46 | 6 | 2976ccv6 | 42405888611 | 42406052476 | 0.0105 | 0.0105 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0317 | 0.6376 | 60.7215 |  |  | 1
47 | 7 | ccb2 | 42405888622 | 42406058680 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0008 | 0 | 0 |  |  | 1
48 | 8 | yb123l | 42405894664 | 42406064884 | 0.0891 | 0.0891 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0757 | 1.5541 | 17.4418 |  |  | 1
49 | 9 | yb123l | 42405894837 | 42406071087 | 0.0923 | 0.0923 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0793 | 1.6285 | 17.644 |  |  | 1
50 | 10 | yb456l | 42405894849 | 42406077291 | 0.0879 | 0.0879 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0723 | 1.4825 | 16.8663 |  |  | 1
51 | 11 | yb456l | 42405894884 | 42406083495 | 0.0884 | 0.0884 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0731 | 1.4999 | 16.9673 |  |  | 1
52 | 12 | yb789l | 42405894895 | 42406089699 | 0.0877 | 0.0877 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0764 | 1.568 | 17.8793 |  |  | 1
53 | 13 | yb789l | 42405894942 | 42406095891 | 0.0886 | 0.0886 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0777 | 1.5958 | 18.0112 |  |  | 1
54 | 14 | sm123l | 42405894953 | 42406102094 | 0.0922 | 0.0922 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0742 | 1.5228 | 16.5165 |  |  | 1
55 | 15 | sm123l | 42405895000 | 42406108298 | 0.0911 | 0.0911 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0733 | 1.5041 | 16.5101 |  |  | 1
56 | 16 | sm456l | 42405895000 | 42406114502 | 0.0992 | 0.0992 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0811 | 1.666 | 16.794 |  |  | 1
57 | 17 | sm456l | 42405895069 | 42406120694 | 0.0821 | 0.0821 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0648 | 1.327 | 16.1635 |  |  | 1
58 | 18 | sm789l | 42405895069 | 42406126898 | 0.0885 | 0.0885 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0911 | 1.8748 | 21.1846 |  |  | 1
59 | 19 | sm789l | 42405895104 | 42406133101 | 0.0861 | 0.0861 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0867 | 1.7826 | 20.7036 |  |  | 1
60 | 20 | 2976ccv5 | 42405907118 | 42406139305 | 0.0146 | 0.0146 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0437 | 0.8875 | 60.7895 |  |  | 1
61 | 21 | 2976ccv7 | 42405907222 | 42406145497 | 0.0142 | 0.0142 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0428 | 0.8688 | 61.1817 |  |  | 1
62 | 22 | ccb3 | 42405907222 | 42406151701 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0003 | 0 | 0 |  |  | 1
63 | 23 | blk | 42406653333 | 42406661793 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0028 | 0.0347 | 0.0347 |  |  | 1
64 | 24 | blk | 42406653368 | 42406668020 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0003 | 0 | 0 |  |  | 1
65 | 25 | dorm-primer3 | 42406653379 | 42406674224 | 0.0502 | 0.0502 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0206 | 20.9068 | 416.4706 |  |  | 1
66 | 26 | dorm-primer4 | 42406653472 | 42406680428 | 0.0529 | 0.0529 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0212 | 21.4774 | 406.0002 |  |  | 1
67 | 27 | blk | 42406653564 | 42406686631 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.002 | 0.0189 | 0.0189 |  |  | 1
68 | 28 | 2976-ccv8 | 42406653587 | 42406692835 | 0.0163 | 0.0163 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0551 | 1.1256 | 69.0557 |  |  | 1
69 | 29 | 2976-ccv9 | 42406653645 | 42406699050 | 0.0152 | 0.0152 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0527 | 1.0744 | 70.6846 |  |  | 1
70 | 30 | ccb4 | 42406653657 | 42406705254 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0002 | 0 | 0 |  |  | 1
71 | 31 | apple-qcs5 | 42406653680 | 42406711458 | 0.0057 | 0.0057 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0114 | 0.2147 | 37.6739 |  |  | 1
72 | 32 | apple-qcs6 | 42406653761 | 42406717685 | 0.0049 | 0.0049 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0124 | 0.2356 | 48.0756 |  |  | 1
73 | 33 | method-blk | 42406653761 | 42406723888 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.001 | 0 | 0 |  |  | 1
74 | 35 | 2976ccv10 | 42406730219 | 42406733136 | 0.0124 | 0.0124 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0424 | 0.8611 | 69.4432 |  |  | 1
75 | 36 | 2976ccv11 | 42406730289 | 42406739386 | 0.0128 | 0.0128 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0408 | 0.8271 | 64.6189 |  |  | 1
76 | 37 | ccb4 | 42406730300 | 42406745590 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
77 | 38 | apple-qcs | 42406730347 | 42406751793 | 0.0054 | 0.0054 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0144 | 0.2772 | 51.3387 |  |  | 1
78 | 39 | apple-qcs2 | 42406764942 | 42406766377 | 0.0047 | 0.0047 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0098 | 0.1808 | 38.4643 |  |  | 1
79 | 40 | blk | 42406765567 | 42406772638 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0004 | 0 | 0 |  |  | 1
80 | 1 | sp123w | 42406769548 | 42406845034 | 0.1278 | 0.1278 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0154 | 0.2981 | 2.3322 |  |  | 1
81 | 2 | sp123w | 42406769606 | 42406851284 | 0.09 | 0.09 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0089 | 0.1619 | 1.7987 |  |  | 1
82 | 3 | sp456w | 42406769606 | 42406857488 | 0.091 | 0.091 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0097 | 0.1793 | 1.9707 |  |  | 1
83 | 4 | sp456w | 42406769652 | 42406863692 | 0.0908 | 0.0908 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0096 | 0.1772 | 1.9521 |  |  | 1
84 | 5 | sp789w | 42406769652 | 42406869895 | 0.0889 | 0.0889 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0096 | 0.1772 | 1.9938 |  |  | 1
85 | 6 | sp789w | 42406769699 | 42406876099 | 0.09 | 0.09 | normal | finished | yy160121.c80 | yy150929.m80 | 0.008 | 0.1439 | 1.5991 |  |  | 1
86 | 7 | fir123w | 42406769699 | 42406882303 | 0.0873 | 0.0873 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0078 | 0.1398 | 1.6009 |  |  | 1
87 | 8 | fir123w | 42406769745 | 42406888495 | 0.0913 | 0.0913 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0077 | 0.1377 | 1.5079 |  |  | 1
88 | 9 | fir456w | 42406769756 | 42406894699 | 0.0882 | 0.0882 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0061 | 0.1043 | 1.1831 |  |  | 1
89 | 10 | fir456w | 42406769791 | 42406900902 | 0.0909 | 0.0909 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0059 | 0.1002 | 1.1021 |  |  | 1
90 | 11 | fir789w | 42406769803 | 42406907106 | 0.0934 | 0.0934 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0068 | 0.1189 | 1.2733 |  |  | 1
91 | 12 | fir789w | 42406769837 | 42406913310 | 0.0929 | 0.0929 | normal | finished | yy160121.c80 | yy150929.m80 | 0.006 | 0.1023 | 1.1008 |  |  | 1
92 | 13 | ab123w | 42406770000 | 42406919502 | 0.0907 | 0.0907 | normal | finished | yy160121.c80 | yy150929.m80 | 0.007 | 0.1231 | 1.3571 |  |  | 1
93 | 14 | ab123w | 42406770046 | 42406925706 | 0.0911 | 0.0911 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0058 | 0.0989 | 1.0859 |  |  | 1
94 | 15 | ab456w | 42406770057 | 42406931909 | 0.0874 | 0.0874 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0106 | 0.1981 | 2.2663 |  |  | 1
95 | 16 | ab456w | 42406770104 | 42406938113 | 0.0888 | 0.0888 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0108 | 0.2022 | 2.2775 |  |  | 1
96 | 17 | ab789w | 42406770104 | 42406944317 | 0.088 | 0.088 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0078 | 0.1398 | 1.5881 |  |  | 1
97 | 18 | ab789w | 42406770150 | 42406950509 | 0.0914 | 0.0914 | normal | finished | yy160121.c80 | yy150929.m80 | 0.009 | 0.1648 | 1.8025 |  |  | 1
98 | 19 | 2976ccv11 | 42406770162 | 42406956712 | 0.0143 | 0.0143 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0504 | 1.0271 | 71.824 |  |  | 1
99 | 20 | 2976ccv12 | 42406770208 | 42406962916 | 0.0156 | 0.0156 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0456 | 0.9271 | 59.4296 |  |  | 1
100 | 21 | ccb5 | 42406770219 | 42406969120 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
101 | 22 | sm123w | 42406770289 | 42406975312 | 0.0867 | 0.0867 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0084 | 0.1523 | 1.7561 |  |  | 1
102 | 23 | sm123w | 42406813125 | 42406981539 | 0.0901 | 0.0901 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0059 | 0.1002 | 1.1119 |  |  | 1
103 | 24 | sm456w | 42406813125 | 42406987743 | 0.0928 | 0.0928 | normal | finished | yy160121.c80 | yy150929.m80 | 0.003 | 0.0398 | 0.4286 |  |  | 1
104 | 25 | sm456w | 42406813159 | 42406993946 | 0.0906 | 0.0906 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0028 | 0.0356 | 0.393 |  |  | 1
105 | 26 | sm789w | 42406813171 | 42407000150 | 0.0874 | 0.0874 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0032 | 0.0439 | 0.5028 |  |  | 1
106 | 27 | sm789w | 42406813206 | 42407006342 | 0.0877 | 0.0877 | normal | finished | yy160121.c80 | yy150929.m80 | 0.004 | 0.0606 | 0.691 |  |  | 1
107 | 28 | yb123w | 42406813217 | 42407012546 | 0.0895 | 0.0895 | normal | finished | yy160121.c80 | yy150929.m80 | 0.009 | 0.1648 | 1.8408 |  |  | 1
108 | 29 | yb123w | 42406813263 | 42407018750 | 0.0867 | 0.0867 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0092 | 0.1689 | 1.9483 |  |  | 1
109 | 30 | yb456w | 42406813275 | 42407024953 | 0.0922 | 0.0922 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0142 | 0.2733 | 2.9642 |  |  | 1
110 | 31 | yb456w | 42406813321 | 42407031145 | 0.0906 | 0.0906 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0129 | 0.246 | 2.7151 |  |  | 1
111 | 32 | yb789w | 42406813321 | 42407037349 | 0.0915 | 0.0915 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0161 | 0.3118 | 3.4079 |  |  | 1
112 | 33 | yb789w | 42406813356 | 42407043553 | 0.089 | 0.089 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0166 | 0.3231 | 3.6298 |  |  | 1
113 | 34 | hb-soil | 42406813368 | 42407049756 | 0.0787 | 0.0787 | normal | finished | yy160121.c80 | yy150929.m80 | 0.002 | 0.0189 | 0.2407 |  |  | 1
114 | 35 | 2976ccv13 | 42406813414 | 42407055949 | 0.0125 | 0.0125 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0389 | 0.7875 | 63.0037 |  |  | 1
115 | 36 | 2976ccv14 | 42406813414 | 42407062152 | 0.0138 | 0.0138 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0432 | 0.8771 | 63.5588 |  |  | 1
116 | 37 | ccb6 | 42406841087 | 42407068356 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0006 | 0 | 0 |  |  | 1
117 | 38 | blk | 42406842129 | 42407074560 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0004 | 0 | 0 |  |  | 1
118 | 39 | blk | 42406842152 | 42407080775 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0004 | 0 | 0 |  |  | 1
119 | 40 | blk | 42406842164 | 42407086990 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
120 | 1 | blk | 42408062696 | 42408063530 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.003 | 0.0407 | 0.0407 |  |  | 1
121 | 2 | blk | 42408063125 | 42408069791 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0002 | 0 | 0 |  |  | 1
122 | 3 | blk | 42408063125 | 42408075995 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0004 | 0 | 0 |  |  | 1
123 | 4 | blk | 42408063136 | 42408082199 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0004 | 0 | 0 |  |  | 1
124 | 5 | blk | 42408063136 | 42408088391 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0001 | 0 | 0 |  |  | 1
125 | 6 | blk | 42408063148 | 42408094606 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0006 | 0 | 0 |  |  | 1
126 | 7 | blk | 42408063148 | 42408100844 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0 | 0 | 0 |  |  | 1
127 | 8 | blk | 42408063159 | 42408107060 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0006 | 0 | 0 |  |  | 1
128 | 9 | blk | 42408063159 | 42408113287 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0 | 0 | 0 |  |  | 1
129 | 10 | blk | 42408063252 | 42408119513 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0004 | 0 | 0 |  |  | 1

## raw data2
LIMS export of <yy160122mdl.d80>
Version 1.0
Date 2016-01-26-15-59-58
Nr | Pos | SampleName | CreationDate | ProcessTime | Weigth [g] | ConcentrateWeight [g] | Type | State | CalFile | MethodFile | Height | Hg [ng] | [ug/kg] | DilutionWeight | FinalVolume | CalFactor | DilutionFactor | Statistics | Group | Remark
1 | 1 | blk | 42391660844 | 42391680740 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0037 | 0 | 0 |  |  | 1
2 | 2 | blk | 42391660891 | 42391686643 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0006 | 0 | 0 |  |  | 1
3 | 3 | 2976ccv1 | 42391661087 | 42391692696 | 0.0186 | 0.0186 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0552 | 1.1983 | 64.426 |  |  | 1
4 | 4 | 2976ccv2 | 42391661307 | 42391698912 | 0.0156 | 0.0156 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0478 | 1.0191 | 65.3258 |  |  | 1
5 | 5 | ccb1 | 42391661319 | 42391705046 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
6 | 6 | applea | 42391661412 | 42391711192 | 0.0048 | 0.0048 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0139 | 0.198 | 41.242 |  |  | 1
7 | 7 | applea | 42391661458 | 42391717476 | 0.0056 | 0.0056 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0138 | 0.1955 | 34.9177 |  |  | 1
8 | 8 | applea | 42391661458 | 42391723634 | 0.005 | 0.005 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0111 | 0.1301 | 26.0281 |  |  | 1
9 | 9 | applea | 42391661469 | 42391729791 | 0.0046 | 0.0046 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0108 | 0.1229 | 26.7117 |  |  | 1
10 | 10 | applea | 42391661469 | 42391736377 | 0.005 | 0.005 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0159 | 0.2464 | 49.281 |  |  | 1
11 | 11 | applea | 42391661481 | 42391743148 | 0.0049 | 0.0049 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0115 | 0.1411 | 28.7923 |  |  | 1
12 | 12 | applea | 42391661481 | 42391749050 | 0.0055 | 0.0055 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0124 | 0.1629 | 29.6145 |  |  | 1
13 | 13 | applea | 42391661516 | 42391755162 | 0.005 | 0.005 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0108 | 0.1241 | 24.8257 |  |  | 1
14 | 14 | applea | 42391661539 | 42391761296 | 0.005 | 0.005 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0117 | 0.1459 | 29.1852 |  |  | 1
15 | 15 | applea | 42391661539 | 42391767430 | 0.0046 | 0.0046 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0118 | 0.1483 | 32.2496 |  |  | 1
16 | 16 | 2976ccv3 | 42391661550 | 42391773541 | 0.0101 | 0.0101 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0295 | 0.577 | 57.1321 |  |  | 1
17 | 17 | 2976ccv4 | 42391661608 | 42391779675 | 0.0151 | 0.0151 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0446 | 0.9427 | 62.4337 |  |  | 1
18 | 18 | ccb2 | 42391661678 | 42391785810 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0006 | 0 | 0 |  |  | 1
19 | 19 | blk | 42391661724 | 42391791979 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0 | 0 | 0 |  |  | 1
20 | 20 | blk | 42391661747 | 42391798263 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
21 | 21 | blk | 42391661759 | 42391804548 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
22 | 22 | blk | 42391661770 | 42391810868 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.001 | 0 | 0 |  |  | 1
23 | 23 | blk | 42391661770 | 42391817175 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0012 | 0 | 0 |  |  | 1
24 | 24 | blk | 42391661793 | 42391823483 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
25 | 25 | blk | 42391661805 | 42391829780 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
26 | 26 | blk | 42391661805 | 42391836087 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
27 | 27 | blk | 42391661817 | 42391842407 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0009 | 0 | 0 |  |  | 1
28 | 28 | blk | 42391661817 | 42391848715 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
29 | 29 | 2976ccv5 | 42391661828 | 42391855011 | 0.0155 | 0.0155 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0443 | 0.9349 | 60.3142 |  |  | 1
30 | 30 | 2976ccv6 | 42391661886 | 42391861319 | 0.0133 | 0.0133 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0411 | 0.858 | 64.5098 |  |  | 1
31 | 31 | ccb3 | 42391661886 | 42391867476 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0006 | 0 | 0 |  |  | 1
32 | 32 | blk | 42392656064 | 42392661469 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0029 | 0 | 0 |  |  | 1
33 | 33 | blk | 42392656099 | 42392667361 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0 | 0 | 0 |  |  | 1
34 | 34 | applea | 42392656111 | 42392673425 | 0.0482 | 0.0482 | normal | finished | yy160121.c80 | yy150929.m80 | 0.1042 | 2.3854 | 49.4901 |  |  | 1
35 | 35 | applea | 42392659537 | 42392679629 | 0.0055 | 0.0055 | normal | finished | yy160121.c80 | yy150929.m80 | 0.011 | 0.129 | 23.4495 |  |  | 1
36 | 36 | applea | 42392686782 | 42392688043 | 0.0052 | 0.0052 | normal | finished | yy160121.c80 | yy150929.m80 | 0.014 | 0.3354 | 64.5021 |  |  | 1
37 | 37 | applea | 42392695185 | 42392696284 | 0.0107 | 0.0107 | normal | finished | yy160121.c80 | yy150929.m80 | 0.022 | 0.3954 | 36.9521 |  |  | 1
38 | 38 | applea | 42392704548 | 42392706458 | 0.0046 | 0.0046 | normal | finished | yy160121.c80 | yy150929.m80 | 0.011 | 0.1261 | 27.4151 |  |  | 1
39 | 39 | applea | 42392704583 | 42392712361 | 0.0105 | 0.0105 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0229 | 0.4137 | 39.4018 |  |  | 1
40 | 40 | applea | 42392716793 | 42392719178 | 0.0054 | 0.0054 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0124 | 0.2356 | 43.6242 |  |  | 1
41 | 1 | applea | 42392718854 | 42392725081 | 0.011 | 0.011 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0236 | 0.4689 | 42.6235 |  |  | 1
42 | 2 | blk | 42392790821 | 42392802002 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.001 | 0 | 0 |  |  | 1
43 | 3 | blk | 42392790856 | 42392807858 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0004 | 0 | 0 |  |  | 1
44 | 4 | applea | 42392790868 | 42392813912 | 0.0048 | 0.0048 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0139 | 0.2668 | 55.5864 |  |  | 1
45 | 5 | applea | 42392790995 | 42392820138 | 0.0054 | 0.0054 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0122 | 0.2314 | 42.8527 |  |  | 1
46 | 6 | applea | 42392790995 | 42392826296 | 0.0054 | 0.0054 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0125 | 0.2377 | 44.0099 |  |  | 1
47 | 7 | applea | 42392791006 | 42392832476 | 0.0047 | 0.0047 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0109 | 0.2039 | 43.3808 |  |  | 1
48 | 8 | applea | 42392791006 | 42392838645 | 0.0049 | 0.0049 | normal | finished | yy160121.c80 | yy150929.m80 | 0.011 | 0.2064 | 42.1244 |  |  | 1
49 | 9 | applea | 42392791018 | 42392844814 | 0.0056 | 0.0056 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0117 | 0.221 | 39.4625 |  |  | 1
50 | 10 | applea | 42392791018 | 42392850972 | 0.0052 | 0.0052 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0117 | 0.221 | 42.4981 |  |  | 1
51 | 11 | applea | 42392791030 | 42392857141 | 0.0054 | 0.0054 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0132 | 0.2522 | 46.71 |  |  | 1
52 | 12 | applea | 42392791030 | 42392863321 | 0.0056 | 0.0056 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0122 | 0.2324 | 41.5016 |  |  | 1
53 | 13 | applea | 42392791041 | 42392869490 | 0.0053 | 0.0053 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0117 | 0.221 | 41.6963 |  |  | 1
54 | 14 | 2976ccv1 | 42392791041 | 42392875659 | 0.0184 | 0.0184 | normal | finished | yy160121.c80 | yy150929.m80 | 0.052 | 1.0604 | 57.6309 |  |  | 1
55 | 15 | 2976ccv7 | 42392791087 | 42392881828 | 0.0174 | 0.0174 | normal | finished | yy160121.c80 | yy150929.m80 | 0.048 | 0.9771 | 56.1547 |  |  | 1
56 | 16 | ccb1 | 42392791087 | 42392887986 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0 | 0 | 0 |  |  | 1
57 | 17 | blk | 42392883668 | 42392927268 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0008 | 0 | 0 |  |  | 1
58 | 18 | 11 | 42392883738 | 42392933206 | 0.0995 | 0.0995 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0023 | 0.0252 | 0.2532 |  |  | 1
59 | 19 | 12 | 42392883796 | 42392945335 | 0.0992 | 0.0992 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0022 | 0.0231 | 0.233 |  |  | 1
60 | 20 | 13 | 42392883807 | 42392951458 | 0.0985 | 0.0985 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0022 | 0.0231 | 0.2346 |  |  | 1
61 | 21 | 21 | 42392883819 | 42392969930 | 0.0926 | 0.0926 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0023 | 0.0252 | 0.2721 |  |  | 1
62 | 22 | 22 | 42392883842 | 42392976111 | 0.0941 | 0.0941 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0023 | 0.0252 | 0.2677 |  |  | 1
63 | 23 | 23 | 42392883842 | 42392982268 | 0.0958 | 0.0958 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0022 | 0.0231 | 0.2413 |  |  | 1
64 | 24 | 31 | 42392883877 | 42393006898 | 0.0902 | 0.0902 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0029 | 0.0377 | 0.4179 |  |  | 1
65 | 25 | 32 | 42392883900 | 42393019224 | 0.0955 | 0.0955 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0027 | 0.0335 | 0.3511 |  |  | 1
66 | 26 | 33 | 42392883912 | 42393031574 | 0.0921 | 0.0921 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0029 | 0.0377 | 0.4093 |  |  | 1
67 | 27 | 2976ccv7 | 42392883935 | 42393043900 | 0.0187 | 0.0187 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0566 | 1.1562 | 61.8302 |  |  | 1
68 | 28 | 2976ccv8 | 42392884039 | 42393050046 | 0.0202 | 0.0202 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0116 | 0.2194 | 10.8619 |  |  | 1
69 | 29 | ccb6 | 42392884050 | 42393055891 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0006 | 0 | 0 |  |  | 1
Nr | Pos | SampleName | CreationDate | ProcessTime | Weigth [g] | ConcentrateWeight [g] | Type | State | CalFile | MethodFile | Height | Hg [ng] | [ug/kg] | DilutionWeight | FinalVolume | CalFactor | DilutionFactor | Statistics | Group | Remark
1 | 1 | blk | 42395666840 | 42395679675 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0015 | 0.0085 | 0.0085 |  |  | 1
2 | 2 | blk | 42395666886 | 42395685937 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
3 | 3 | dorm-primer1 | 42395666886 | 42395692141 | 0.0558 | 0.0558 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0214 | 21.6676 | 388.3084 |  |  | 1
4 | 4 | dorm-primer2 | 42395666944 | 42395698344 | 0.055 | 0.055 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0221 | 22.3333 | 406.0599 |  |  | 1
5 | 5 | blk | 42395666956 | 42395704548 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0015 | 0.0085 | 0.0085 |  |  | 1
6 | 6 | 2976ccv1 | 42395666967 | 42395710740 | 0.0129 | 0.0129 | normal | finished | yy160121.c80 | yy150929.m80 | 0.038 | 0.7688 | 59.5969 |  |  | 1
7 | 7 | 2976ccv2 | 42395667025 | 42395716944 | 0.0159 | 0.0159 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0502 | 1.0229 | 64.3344 |  |  | 1
8 | 8 | ccb1 | 42395667025 | 42395723148 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
9 | 9 | apple-qcs1 | 42395667048 | 42395729351 | 0.0056 | 0.0056 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0128 | 0.2439 | 43.554 |  |  | 1
10 | 10 | apple-qcs2 | 42395667129 | 42395735555 | 0.0054 | 0.0054 | normal | finished | yy160121.c80 | yy150929.m80 | 0.012 | 0.2272 | 42.0813 |  |  | 1
11 | 11 | method-blk | 42395667129 | 42395741759 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.003 | 0.0398 | 0.0398 |  |  | 1
12 | 12 | applea | 42395741076 | 42395747962 | 0.0055 | 0.0055 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0105 | 0.196 | 35.6355 |  |  | 1
13 | 13 | applea | 42395741620 | 42395754155 | 0.005 | 0.005 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0122 | 0.2314 | 46.281 |  |  | 1
14 | 14 | applea | 42395741620 | 42395760358 | 0.0048 | 0.0048 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0096 | 0.1775 | 36.9764 |  |  | 1
15 | 15 | applea | 42395741620 | 42395766562 | 0.0052 | 0.0052 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0112 | 0.2106 | 40.4953 |  |  | 1
16 | 16 | applea | 42395741631 | 42395772766 | 0.0054 | 0.0054 | normal | finished | yy160121.c80 | yy150929.m80 | 0.012 | 0.2272 | 42.0813 |  |  | 1
17 | 17 | applea | 42395741631 | 42395778969 | 0.005 | 0.005 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0113 | 0.2127 | 42.5317 |  |  | 1
18 | 18 | applea | 42395741643 | 42395785173 | 0.0048 | 0.0048 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0114 | 0.2147 | 44.7378 |  |  | 1
19 | 19 | blk | 42395742662 | 42395791377 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
20 | 20 | 41 | 42395786655 | 42395816099 | 0.0956 | 0.0956 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0038 | 0.0564 | 0.5904 |  |  | 1
21 | 21 | 42 | 42395786793 | 42395828564 | 0.0886 | 0.0886 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0055 | 0.0909 | 1.0261 |  |  | 1
22 | 22 | 43 | 42395786805 | 42395840960 | 0.0937 | 0.0937 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0034 | 0.0481 | 0.5134 |  |  | 1
23 | 23 | 51 | 42395787002 | 42395853368 | 0.0878 | 0.0878 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0031 | 0.0419 | 0.4767 |  |  | 1
24 | 24 | 52 | 42395787025 | 42395865775 | 0.0788 | 0.0788 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0059 | 0.1002 | 1.2713 |  |  | 1
25 | 25 | 53 | 42395787037 | 42395871979 | 0.1057 | 0.1057 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0053 | 0.0877 | 0.8295 |  |  | 1
26 | 26 | 2976ccv3 | 42395787037 | 42395878182 | 0.0202 | 0.0202 | normal | finished | yy160121.c80 | yy150929.m80 | 0.049 | 0.9979 | 49.4021 |  |  | 1
27 | 27 | 2976ccv4 | 42395787037 | 42395884386 | 0.0127 | 0.0127 | normal | finished | yy160121.c80 | yy150929.m80 | 0.04 | 0.8105 | 63.8156 |  |  | 1
28 | 28 | ccb2 | 42395787303 | 42395890578 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0001 | 0 | 0 |  |  | 1
29 | 29 | 61 | 42395787349 | 42395896782 | 0.0967 | 0.0967 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0024 | 0.0273 | 0.2826 |  |  | 1
30 | 30 | 62 | 42395787731 | 42395910636 | 0.097 | 0.097 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0022 | 0.0231 | 0.2383 |  |  | 1
31 | 31 | 63 | 42395787743 | 42395916898 | 0.0888 | 0.0888 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0022 | 0.0231 | 0.2603 |  |  | 1
32 | 32 | 71 | 42395787743 | 42395923101 | 0.1129 | 0.1129 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0025 | 0.0294 | 0.2601 |  |  | 1
33 | 33 | 72 | 42395787800 | 42395935509 | 0.1047 | 0.1047 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0026 | 0.0322 | 0.3073 |  |  | 1
34 | 34 | 73 | 42395787997 | 42395947905 | 0.0968 | 0.0968 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0037 | 0.0544 | 0.5615 |  |  | 1
35 | 35 | 2976ccv5 | 42395787997 | 42395954108 | 0.0127 | 0.0127 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0445 | 0.9042 | 71.1961 |  |  | 1
36 | 36 | 2976ccv6 | 42395788009 | 42395960312 | 0.018 | 0.018 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0594 | 1.2145 | 67.4748 |  |  | 1
37 | 37 | ccb2 | 42395788009 | 42395966516 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
38 | 38 | 81 | 42395890810 | 42395972719 | 0.1104 | 0.1104 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0022 | 0.0231 | 0.2093 |  |  | 1
39 | 39 | 82 | 42395892094 | 42395978923 | 0.1024 | 0.1024 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0021 | 0.0201 | 0.1958 |  |  | 1
40 | 40 | 83 | 42395892106 | 42395985127 | 0.0966 | 0.0966 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0015 | 0.0085 | 0.0883 |  |  | 1
41 | 1 | 91 | 42395899722 | 42396003738 | 0.0973 | 0.0973 | normal | finished | yy160121.c80 | yy150929.m80 | 0.2356 | 225.3681 | 2316.2185 |  |  | 1
42 | 2 | 92 | 42395899803 | 42396009942 | 0.0959 | 0.0959 | normal | finished | yy160121.c80 | yy150929.m80 | 0.2155 | 206.2533 | 2150.7124 |  |  | 1
43 | 3 | 93 | 42395902453 | 42396028541 | 0.1085 | 0.1085 | normal | finished | yy160121.c80 | yy150929.m80 | 0.219 | 209.5818 | 1931.6292 |  |  | 1
44 | 4 | 2976ccv6 | 42395905104 | 42396034745 | 0.0138 | 0.0138 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0572 | 1.169 | 84.7112 |  |  | 1
45 | 5 | 2976ccv7 | 42395905162 | 42396040949 | 0.0113 | 0.0113 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0414 | 0.8402 | 74.3516 |  |  | 1
46 | 6 | ccb4 | 42395905324 | 42396047152 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0016 | 0.0115 | 0.0115 |  |  | 1
47 | 7 | blk | 42395909895 | 42396053344 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0022 | 0.0231 | 0.0231 |  |  | 1
48 | 8 | blk | 42395909930 | 42396059548 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0014 | 0.0064 | 0.0064 |  |  | 1
49 | 9 | blk | 42395909942 | 42396065787 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0015 | 0.0085 | 0.0085 |  |  | 1
50 | 10 | blk | 42395909965 | 42396072037 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0016 | 0.0097 | 0.0097 |  |  | 1
51 | 11 | blk | 42395909965 | 42396078263 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.001 | 0 | 0 |  |  | 1
52 | 12 | blk | 42395909976 | 42396084513 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0014 | 0.0064 | 0.0064 |  |  | 1
53 | 13 | blk | 42395909988 | 42396090752 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0009 | 0 | 0 |  |  | 1
54 | 14 | blk | 42395909988 | 42396097002 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0012 | 0.0018 | 0.0018 |  |  | 1
55 | 15 | blk | 42395909988 | 42396103263 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
56 | 16 | blk | 42396080775 | 42396109490 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0 | 0 | 0 |  |  | 1
57 | 17 | blk | 42396080787 | 42396115740 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
58 | 18 | blk | 42396080787 | 42396121979 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1
59 | 19 | blk | 42396080844 | 42396128240 | 1 | 1 | normal | finished | yy160121.c80 | yy150929.m80 | 0.0007 | 0 | 0 |  |  | 1